PROCESS ANALYTICAL TECHNOLOGY (PAT) IN TEVA API DIVISION

USING NEAR IR (NIR) FOR QUANTITATIVE IMPURITY PROFILE INSTEAD OF HPLC FOR PROCESS UNDERSTANDING

Ilana Nachum, Galina Volter, Adi Kleiman, Moshe Nulman, Ronen Sharf, Revital Rapaport, Amira Rottman

BU Israel API division TEVA, Netanya

Near infra red (NIR) analysis, used as a process analytical technology (PAT) either on-line or in-line, has become a very powerful tool in the pharmaceutical production facilities for the release of intermediates materials.  It has become a necessity for the pharmaceutical and API companies from the FDA perspective, due to the fact that current methods have become limited and less effective.

A NIR on line/in line system was installed in the distillation and circulation pipeline at a production facility of product AA of TEVA API Israel. The purpose was to determine the distillate quality during distillation in order to be able to "correct" the quality on the fly, minimizing out of specification (OOS) results.  Secondary target was the release of batches on-site, at the production facility, avoiding the transfer of samples into the laboratory for analysis.

The Near IR method which is used for routine release of batches in production facility, replaces the HPLC method used in the laboratory, which considered as a specific, precise, accurate and very indicative method. The equivalency between two methods was proven in the validation, eliminating the uncertainty regarding the prospect of  NIR replacing the HPLC method.

NIR Chemometric models were used to identify and quantify impurities of the final intermediate A, using qualitative and quantitative models.  Two related compounds were quantified, impurities A1 and A2.  Models were constructed, using the "Unscrambler" Chemometric software for multivariate data analysis.  After obtaining the spectral data, first derivative was used in order to improve resolution and decreasing offset of baseline, then smoothing was performed by using Savitzky Golay algorithm in order to decrease noise.  The principal component analysis (PCA) and classification by SIMCA were used to create the qualitative model; while the partial least squares (PLS1) method was used to form the quantitative models.

The specificity study was the most important study in the validation of the in-line system, which is usually a closed system.  It was demonstrated that our models can distinguish between different stages in the process, for example, distillation, then circulation.  Also, the possibility to distinguish rejected batches from good batches, as well as the ability to detect situation of "no flow" in the system was proven.

The validated analytical method found to meet the acceptance criteria for specificity, accuracy, precision and robustness.