A NOVEL APPROACH FOR ELECTRON IONIZATION LC-MS WITH SUPERSONIC MOLECULAR BEAMS

Tal Alon, Kfir Gil, Alexander B. Fialkov & Aviv Amirav

School of Chemistry, Tel Aviv University, Tel Aviv 69978, Israel

amirav@tau.ac.il

Electron ionization can significantly benefit LC-MS through the provision of automated library identification and extensive fragment information which is ideally suitable for LC-MS identification of unknown compounds. Thus, bringing back EI to LC-MS is highly valuable if a reliable and robust EI interface can be developed.

We developed a novel EI-LC-MS approach, based on interfacing LC and MS with supersonic molecular beams (SMB) and electron impact ionization of the sample as vibrationally cold compounds in the SMB. The output of a LC was vaporized behind a supersonic nozzle at about 0.1 Bar to suppress cluster formation yet to obtain efficient vibrational cooling, and the sample compounds expended from a 300 µm supersonic nozzle into the vacuum system. Sample vaporization was based on thermally assisted spray formation followed by fast, thermal vaporization of the sample compounds prior to their expansion from the supersonic nozzle. However, the step of thermally assisted spray led to frequent clogging of the solvent delivery tube whose replacement required venting of the MS vacuum system.

Recently, we developed a new approach for EI-LC-MS named Capillary Separated Vaporization Chamber and Nozzle system (CSVCN). We found that the separation of the vaporization chamber and nozzle by a flow restriction capillary is highly beneficial. Accordingly, effective pneumatic spray based sample thermal vaporization occurs at about 2 Atm, while the supersonic expansion is retained at about 0.1 Bar as needed to suppress cluster formation with the solvent vapor, yet to obtain good vibrational cooling. Thus, the CSVCN provides a major improvement of the interface robustness and serviceability, which is now similar to APCI.

The presentation will describe the new CSVCN concept and its many advantages including the ability to mount the vaporization chamber on a GC and have GC-MS and LC-MS in a one system with fast automated interchange between these two modes.