LC/MS/MS QUANTITATVE ANALYSIS OF VITAMIN D: HIDDEN ANALYTICAL ERRORS

Eduard Rogatsky and Daniel T. Stein

Analytical Core Lab, Institute for Clinical and Translational Research, Albert Einstein College of Medicine of Yeshiva University, Bronx, NY USA

eduard.rogatsky@einstein.yu.edu

Vitamin D deficiency is a widespread clinical problem and has been associated with many adverse health outcomes. Analysis of Vitamin D2 (ergocalciferol) and D3 (cholecalciferol) and their major metabolites 25(OH)D2 and 25(OH)D3 has become a high priority topic in clinical analysis. Currently a variety of LC/MS methods have been developed to support vitamin D analysis. These LC/MS methods utilize different transitions, ionization modes, sample preparation strategies, mobile phases and columns.

It should be emphasized that maintaining the maximum S/N and recovery is essential for analysis of vitamin D species. Therefore, we focused on optimization and control of the efficiency of the MRM transitions [M+H+]→[M+H+-18], since the dehydrated vitamin D species were found as dominant by-products for both 25(OH)D2 and 25(OH)D3. Since water loss is a favored process, in-source self –dehydration of vitamin D can decrease S/N. Such an undesirable process was tightly controlled by monitoring of transitions m/z 395.2→377.2 for 25(OH)D2 and m/z 383.2→377.2 for 25(OH)D3 during our method development in addition to “analytical” transitions of m/z. 413.2→395.2 and m/z 401.2→383.2

In-source water loss was found to be insignificant (~8%) within the assay linear range using a methanol/water mobile phase under our experimental conditions. In contrast, if acetonitrile/water was used, severe in-source water loss was observed and ions m/z 395.2 and m/z 383.2 resulting from analyte in-source dehydration, became predominant. This explains the advantage of APCI for vitamin D analysis using acetonitrile instead of methanol, if the dehydrated parent ion is chosen as the precursor for further MRM. A similar strategy has previously been reported for LC/MS analysis of catecholamines.

When acetonitrile was used as mobile phase, and if only the transitions m/z 395.2→377.2 for 25(OH)D2 and m/z 383.2→377.2 for 25(OH)D3 were used for quantitation, the assay standard curve was still acceptably linear; however the area ratio between analyte and its in-source water loss product was non-linear at different analyte concentrations. The concentration-dependent water loss varied by 20%, resulting in a significant quantitation error in this provisionally valid assay. For an acetonitrile-based mobile phase, both transitions (for analyte and for its in-source dehydrated product) should be summed to maintain true full range assay linearity.