Quantification of Free and Metal-Complexed Cyanide

by Tetrathionate Derivatization

Alexey Kamyshny, Jr.1,2, Harry Oduro1, and James Farquhar`

1 - Department of Geology and ESSIC, University of Maryland, College Park, Maryland, USA

2 - Department of Geological & Environmental Studies, The Faculty of Natural Sciences, Ben-Gurion University of the Negev, Beer Sheva, Israel

alexey93@gmail.com

A sensitive and robust method for detection of free and metal-complexed cyanide in solutions is described. The method does not require a distillation step and is applicable for both low ionic strength and sea-water samples. The method is based on the reaction of cyanide with potassium tetrathionate followed by HPLC separation and UV detection of formed thiocyanate. The detection limit of the method is 250 nmol l-1 cyanide (6.5 µg l-1 CN-) without a pre-concentration step. Storage for three days does not significantly change the results. The sum of free and weak metal-complexed cyanide can be measured by tetrathionate derivatization at a pH of 10. The sum of free, weak metal-complexed cyanide, iron(II) and iron(III)-complexed cyanides may be measured by tetrathionate derivatization at pH 4.5. Weighted mean recoveries for free, iron(II), iron(III), nickel(II) and silver(I) complexed cyanide were in the range of 87 – 112% and weighted standard deviations were in the range of 2 – 11%. The method is not applicable for cyanide complexes of gold and cobalt. We illustrate an application of cyanide quantification using pore waters from the Delaware Great Marsh.